LS-APGD Control System

Raspberry Pi Setup & Fresh Install Recipe

GAA Custom Electronics LLC  |  Software Rev 2  |  March 2026

1. Hardware Overview

The LS-APGD control system runs on a Raspberry Pi 3B+ with a custom HAT (R3.0). The HAT provides I2C DAC/ADC, RS-232 level shifting for the Alicat BASIS 2 gas flow controller, USB for the syringe pump, GPIO relay and PWM for the power supply and igniter, and a shutdown GPIO interface.

| Component | Details |
| --- | --- |
| Pi model | Raspberry Pi 3B+ |
| Display | Pi Display 2 (DSI, 720x1280 native, scaled 853x480 landscape) |
| Touch | FT3168 capacitive (Goodix driver, I2C) |
| DAC/ADC | AD5593R at I2C 0x11 (smbus2) |
| ADC (ground ref) | ADS1115 via CircuitPython (main thread only) |
| Gas flow controller | Alicat BASIS 2, /dev/ttyAMA0, 38400 baud, RS-232 |
| Syringe pump | OEM pump, /dev/ttyUSB0, 38400 baud |
| PWM (igniter) | pigpio, GPIO 18, 1 kHz |
| Power relay | GPIO 4 |
| Shutdown pin | GPIO 24 (HIGH = shutdown requested) |

2. OS Installation

Flash a fresh SD card using Raspberry Pi Imager. Use the 64-bit Bookworm Lite image.

## 2.1 Imager Settings

- OS: Raspberry Pi OS (64-bit) — Bookworm

- Hostname: CAMS

- Username: cams

- Enable SSH with password authentication

- Set locale/timezone as required

3. First Boot — System Configuration

SSH into the Pi after first boot and run the following steps in order.

## 3.1 Update System

## sudo apt update && sudo apt upgrade -y

## 3.2 Install apt Packages

## sudo apt install -y python3-tk python3-pigpio python3-smbus2 \

## python3-serial python3-rpi.gpio wlr-randr git imagemagick

## 3.3 Install pip Packages

## pip3 install adafruit-blinka --break-system-packages

## pip3 install adafruit-circuitpython-ads1x15 --break-system-packages

## 3.4 Build pigpio from Source

## pigpio is not in the Bookworm apt repos — build and install manually:

## git clone https://github.com/joan2937/pigpio.git

## cd pigpio && make && sudo make install

## Create the systemd service file:

## sudo nano /etc/systemd/system/pigpiod.service

## Paste the following content:

## [Unit]

## Description=Pigpio daemon

## After=network.target

## [Service]

## Type=forking

## ExecStart=/usr/local/bin/pigpiod

## ExecStop=/bin/systemctl kill pigpiod

## Restart=on-failure

## [Install]

## WantedBy=multi-user.target

## sudo systemctl daemon-reload

## sudo systemctl enable pigpiod

## sudo systemctl start pigpiod

## 3.5 User Groups and Permissions

## sudo usermod -a -G dialout,gpio,i2c,spi cams

## Add udev rule so ttyAMA0 is accessible without root:

## sudo nano /etc/udev/rules.d/99-ttyAMA0.rules

## File content:

## KERNEL=="ttyAMA0", GROUP="dialout", MODE="0660"

## sudo udevadm control --reload-rules && sudo udevadm trigger

## 3.6 Disable Serial Console on ttyAMA0

## The serial-getty service grabs ttyAMA0 and blocks the gas flow controller. Disable and mask it permanently:

## sudo systemctl stop serial-getty@ttyAMA0.service

## sudo systemctl disable serial-getty@ttyAMA0.service

## sudo systemctl mask serial-getty@ttyAMA0.service

## NOTE: Using 'disable' alone is not enough on Bookworm — 'mask' is required to prevent systemd from ever respawning it.

4. Boot Configuration

## 4.1 /boot/firmware/config.txt

## Edit with: sudo nano /boot/firmware/config.txt

## Ensure these lines are present. Do NOT add dtoverlay=vc4-kms-dsi-ili9881-7inch — it kills the display.

## dtparam=i2c_arm=on

## dtparam=spi=on

## dtparam=audio=on

## camera_auto_detect=1

## display_auto_detect=0

## auto_initramfs=1

## dtoverlay=vc4-kms-v3d

## max_framebuffers=2

## disable_fw_kms_setup=1

## arm_64bit=1

## disable_overscan=1

## arm_boost=1

## [all]

## enable_uart=1

## dtoverlay=disable-bt

## 4.2 /boot/firmware/cmdline.txt

## Verify this file looks like the following (all on one line). Do NOT add rd.systemd.show_status=false or systemd.show_status=false — this breaks the display and SSH:

## console=serial0,115200 console=tty1 root=PARTUUID=xxxx-xx rootfstype=ext4

## fsck.repair=yes rootwait quiet splash plymouth.ignore-serial-consoles

## loglevel=0 vt.global_cursor_default=0 cfg80211.ieee80211_regdom=US

5. Display Configuration

## 5.1 Set Display Scale

## The Pi Display 2 runs at 720x1280 native (portrait). Scale to 1.5 for a usable landscape resolution of 853x480 via raspi-config or Raspberry Pi Config GUI. Verify:

## cat ~/.config/kanshi/config

## Should contain:

## profile {

## output DSI-1 enable scale 1.500000 mode 720x1280@60.038 position 0,0 transform 90

## }

## 5.2 Touch Mapping (labwc)

## mkdir -p ~/.config/labwc

## nano ~/.config/labwc/rc.xml

## Content:

## <?xml version="1.0"?>

## <openbox_config xmlns="http://openbox.org/3.4/rc">

## <touch deviceName="10-005d Goodix Capacitive TouchScreen"

## mapToOutput="DSI-1" mouseEmulation="yes"/>

## </openbox_config>

## 5.3 App Autostart

## nano ~/.config/labwc/autostart

## Content:

## sleep 2

## cd /home/cams/LS_APGD && python3 main.py &

6. Boot Splash Screen

Set up the logo as the Plymouth boot splash. Resize to 720x1280 (native resolution — Plymouth runs before display scaling):

convert ~/splash.png -resize 720x1280 -background black \

-gravity center -extent 720x1280 splash_native.png

sudo cp splash_native.png /usr/share/plymouth/themes/pix/splash.png

sudo update-initramfs -u

7. Application Files — Deploy from GitHub

The LS-APGD application is hosted on GitHub. Clone the repo directly to the Pi instead of copying files manually:

## 7.1 Clone the Repository

## git clone https://github.com/anderson509chris/LS_APGD.git /home/cams/LS_APGD

## This creates /home/cams/LS_APGD/ with all 7 application modules:

- main.py — entry point, hardware init, serial connections

- hardware.py — GPIO, pigpio PWM, AD5593R I2C DAC/ADC, ADS1115, SensorPoller thread

- pump.py — syringe pump serial (/dev/ttyUSB0, 38400 baud)

- gas.py — Alicat BASIS 2 flow controller (/dev/ttyAMA0, 38400 baud, BASIS 2 protocol)

- settings_manager.py — read/write LSAPGD_SAVE.txt

- dialogs.py — number pad dialog, show_message

- gui_pages.py — full App class, all 6 tabs, sensor display loop

## 7.2 Updating the App

## To pull the latest version of the app from GitHub:

## cd /home/cams/LS_APGD && git pull origin main

## 7.4 Settings File

## The app reads and writes /home/cams/LS_APGD/LSAPGD_SAVE.txt on startup and shutdown. This file is excluded from git (.gitignore) so it is never overwritten by a pull. If the file does not exist the app creates it with defaults on first run.

8. Hardware Notes

## 8.1 Alicat BASIS 2 Gas Flow Controller

## The BASIS 2 communicates via RS-232 through the HAT's MAX3232 level shifter. The HAT's RS-232 option jumpers must be installed. The JST-GH 6-pin connector pinout is:

| Pin | Function |
| --- | --- |
| 1 | Analog setpoint input (not used — serial control) |
| 2 | Analog signal output (not used) |
| 3 | Ground |
| 4 | Power in 12-24V |
| 5 | RS-232 Rx (receives commands from Pi) |
| 6 | RS-232 Tx (sends responses to Pi) |

## The app uses BASIS 2 ASCII Protocol 2. Key commands:

- Poll device: A\r — returns full data frame

- Set setpoint source: ALSS u\r — unsaved digital mode

- Set flow: AS 100.000\r — sets flow to 100 mL/min

- Stop flow: AS 0\r

- Set gas: AGS 7\r — 7 = Helium

NOTE: The BASIS 2 requires up to 1.0 second timeout to respond. The serial-getty@ttyAMA0 service must be masked or it will block the port.

## 8.2 GPIO Pin Summary

| GPIO | Function |
| --- | --- |
| GPIO 2/3 | I2C SDA/SCL — AD5593R + ADS1115 |
| GPIO 4 | Power supply relay output |
| GPIO 14/15 | UART TX/RX — gas flow controller (ttyAMA0) |
| GPIO 18 | Igniter PWM (pigpio, 1 kHz) |
| GPIO 24 | Shutdown pin output (HIGH = shutdown requested) |

## 8.3 ADS1115 Thread Safety

## The ADS1115 uses the Adafruit CircuitPython I2C stack which is not thread-safe. It must only be read from the main thread. The SensorPoller background thread reads only the AD5593R via smbus2. The ADS1115 ground reference is read directly on the main thread in _update_voltage_display() when power is OFF.

9. Quick Checklist

Use this checklist when building a new SD card from scratch:

- Flash Raspberry Pi OS Bookworm 64-bit — hostname=CAMS, user=cams, enable SSH

- sudo apt update && sudo apt upgrade -y

- Install apt packages (Section 3.2)

- Install pip packages (Section 3.3)

- Build pigpio from source and create pigpiod.service (Section 3.4)

- sudo usermod -a -G dialout,gpio,i2c,spi cams

- Add udev rule for ttyAMA0 (Section 3.5)

- Stop, disable, and MASK serial-getty@ttyAMA0.service (Section 3.6)

- Edit /boot/firmware/config.txt (Section 4.1)

- Verify /boot/firmware/cmdline.txt (Section 4.2)

- Set display scale to 1.5 via raspi-config or Raspberry Pi Config GUI

- Verify ~/.config/kanshi/config (Section 5.1)

- Create ~/.config/labwc/rc.xml touch mapping (Section 5.2)

- Create ~/.config/labwc/autostart (Section 5.3)

- Clone app from GitHub: git clone https://github.com/anderson509chris/LS_APGD.git /home/cams/LS_APGD

- Set up Plymouth splash image (Section 6)

- Reboot and verify app starts, touch works, gas valve responds

- Create SD card backup image (Section 9)

GAA Custom Electronics LLC — LS-APGD Setup Recipe — Software Rev 2